[image: image1.jpg]


Vermont Department of Health Laboratory

195 Colchester Avenue, P.O. Box 1125, Burlington VT  05402-1125

DataMaster Technical Services (ph) 802-863-7641; (fax) 802-863-7632


DATAMASTER TECHNICAL SUPPORT INQUIRY

Agency/Site: VDHL

DataMaster Serial #: 100158
Date: 11
/16
/2011

Requestor: Kimball

Instrument Supervisor: 

Phone #: 

Taken By: HARNOIS

Problem: Filter Wheel Pin shifted, Higher then expected Xq. 



Laboratory Response:




Work Performed:

11-16-2011 – Found that Filter Wheel had shifted in the solenoid, replaced filter wheel solenoid with one from stock, reassembled adjusted lamp and cooler to get to 0.000 V on the detector (+/- 0.100) tested with no filter wheel errors. Instrument to be Calibration. SH

 FORMCHECKBOX 
 On Site
 FORMCHECKBOX 
 In-house



Attachments: ____0__ pages
Performed By: Steven Harnois

Date:
11/16/2011
Final Review: Kirk L. Kimball

Date:      
11/30/2011

ALC 626 Rev 3 06/11 

Printed on: November 30, 2011  13:55

